Digital Imaging and Communications in Medicine (DICOM)

Supplement 237: High-Resolution ECG

*Prepared by: Working Group 32*

**DICOM Standards Committee, Working Group 6**

1300 N. 17th Street, Suite 900

Rosslyn, Virginia 22209 USA

Status: Version 4, Public Comment Dec. 2022

Developed pursuant to DICOM Work Item **2018-09-D**

**Table of Contents**

Document History 4

Open Issues 4

Closed Issues 4

Scope and Field of Application 4

Changes to NEMA Standards Publications PS3.2 Digital Imaging and Communications in Medicine (DICOM) Part 2: Conformance 5

Changes to NEMA Standards Publications PS3.3 Digital Imaging and Communications in Medicine (DICOM) Part 3: Information Object Definitions 5

A.34.4.1 General ECG IOD Description 6

A.34.xx High Resolution ECG IOD 6

A.34.xx.1 High Resolution ECG IOD Description 6

A.34.xx.2 High Resolution ECG IOD Entity-Relationship Model 6

A.34.xx.3 High Resolution ECG IOD Module Table 6

A.34.xx.4 High-Resolution ECG IOD Content Constraints 7

A.34.xx.4.1 Modality 7

A.34.xx.4.2 Waveform Sequence 7

A.34.xx.4.3 Number of Waveform Channels 7

A.34.xx.4.4 Sampling Frequency 7

A.34.xx.4.5 Channel Source and Channel Source Modifiers 7

A.34.xx.4.6 Waveform Sample Interpretation 7

A.34.xx.4.7 Waveform Annotation Module 7

Changes to NEMA Standards Publications PS3.4 Digital Imaging and Communications in Medicine (DICOM) Part 4: Service Class Specifications 8

B.5 Standard SOP classes 8

Changes to NEMA Standards Publications PS3.6 Digital Imaging and Communications in Medicine (DICOM) Part 6: Data Dictionary 8

Changes to NEMA Standards Publications PS 3.17 Digital Imaging and Communications in Medicine (DICOM) Part 17: Explanatory Information 9

SSSS.1.5.1 Mapping of Polysomnographic Data to DICOM 9

# Document History

|  |  |  |  |
| --- | --- | --- | --- |
| 2022/09/02 | Version 0 |  | Initial version, fragmentary |
| 2022/10/01 | Version 1 |  | after First Read, recommendations incorporated |
| 2022/11/18 | Version 3 |  | pre-Public Comment |
| 2022/11/21 | Version 4 |  | Public Comment |

# Open Issues

|  |  |
| --- | --- |
|  |  |

# Closed Issues

|  |  |
| --- | --- |
| 1. | Q: Should there be a general recommendation to use the new High Resolution ECG SOP Class, or should both be possible (dependent on the properties of the recording device)?WG-32 is considering to recommend usage of High Resolution ECG for all neurophysiology studies.Feedback from implementations of ECG viewers and storage systems is encouraged.A: Depending on the capabilities of the recording device, the vendor decides which SOP class to support. |
| 2. | Q: Should the Acquisition Context contain a reference to the related EEG SOP Instance and vice versa, if the ECG is acquired in a neurophysiology study? A: No, SOP Instances in a study may be created by different equipment, which don't know of each other. The presence of a reference cannot be guaranteed |

# Scope and Field of Application

This supplement defines a new ECG Waveform SOP Class (based on the existing General ECG SOP Class) with fewer constraints. The General ECG SOP class can store waveform with 16 bits per sample. This High Resolution SOP class permits 32 bits per sample as well.

In clinical neurophysiology it is common practice to acquire ECG data together with the routine scalp EEG or in case of a sleep study.

# Changes to NEMA Standards Publications PS3.2Digital Imaging and Communications in Medicine (DICOM)Part 2: Conformance

Add the new SOP Classes to Table A.1-2. UID Values

Table A.1-2. UID Values

|  |  |  |
| --- | --- | --- |
| UID Value​  | UID Name​  | Category​ |
| …. ​  |  |  |
| **1.2.840.10008.5.1.4.1.1.9.1.x** | **High Resolution ECG Waveform Storage SOP Class** | **Transfer** |
| … |  |  |

# **Changes to NEMA Standards Publications PS3.3****Digital Imaging and Communications in Medicine (DICOM)****Part 3: Information Object Definitions**

Add new IODs to Overview Table PS3.3 Table A.1-8:

Table A.1-8. Composite Information Object Modules Overview – Waveforms

|  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| IOD Modules | … | General ECG WF | … | High Resolution ECG WF | … |
| Patient |  | M |  | **M** |  |
| Clinical Trial Subject |  | U |  | **U** |  |
| General Study |  | M |  | **M** |  |
| Patient Study |  | U |  | **U** |  |
| Clinical Trial Study |  | U |  | **U** |  |
| General Series |  | M |  | **M** |  |
| Clinical Trial Series |  | U |  | **U** |  |
| Synchronization |  | U |  | **U** |  |
| General Equipment |  | M |  | **M** |  |
| Enhanced General Equipment |  |  |  | **M** |  |
| Waveform Identification |  | M |  | **M** |  |
| Waveform |  | M |  | **M** |  |
| Waveform Annotation |  | C |  | **C** |  |
| Acquisition Context |  | M |  | **M** |  |
| SOP Common |  | M |  | **M** |  |

Extend A.34.4.1 General ECG IOD Description

### A.34.4.1 General ECG IOD Description

The General ECG IOD is the specification of digitized electrical signals from the patient cardiac conduction system collected on the body surface, which has been acquired by an ECG modality or by an ECG acquisition function within an imaging modality **or by another recording device**.

Add the following new content to PS3.3Section A.34.xx

## A.34.xx High Resolution ECG IOD

### A.34.xx.1 High Resolution ECG IOD Description

The High Resolution ECG IOD is the specification of digitized electrical signals from the patient cardiac conduction system collected on the body surface, which has been acquired by an ECG modality or by an ECG acquisition function within an imaging modality or by another recording device.

Note

This IOD differs from the General ECG IOD by allowing higher sampling frequencies and greater bit depth.

### A.34.xx.2 High Resolution ECG IOD Entity-Relationship Model

This IOD uses the E-R Model in Section A.1-2, with only the Waveform IE below the Series IE.

### A.34.xx.3 High Resolution ECG IOD Module Table

Table A.34.xx.1-1 specifies the Modules of the High Resolution ECG IOD

Table A.34.xx.1-1- High Resolution ECG IOD Modules

|  |  |  |  |
| --- | --- | --- | --- |
| IE | Module | Reference | Usage |
| Patient | Patient | C.7.1.1 | M |
| Clinical Trial Subject | C.7.1.3 | U |
| Study | General Study | C.7.2.1 | M |
| Patient Study | C.7.2.2 | U |
| Clinical Trial Study | C.7.2.3 | U |
| Series | General Series | C.7.3.1 | M |
| Clinical Trial Series | C.7.3.2 | U |
| Presentation Series | C.7.3.3 | M |
| Frame of Reference | Synchronization | C.7.4.2 | C – Required if time synchronization was applied. |
| Equipment | General Equipment | C.7.5.1 | M |
| Enhanced General Equipment | C.7.5.2 | M |
| Waveform  | Waveform Identification | C.10.8 | M |
| Waveform | C.10.9 | M |
| Acquisition Context | C.7.6.14 | M |
| Waveform Annotation | C.10.10 | C- Required if annotation is present. |
| SOP Common | C.12.1 | M |

### A.34.xx.4 High Resolution ECG IOD Content Constraints

#### A.34.xx.4.1 Modality

The value of Modality (0008,0060) shall be ECG.

#### A.34.xx.4.2 Waveform Sequence

The number of Waveform Sequence (5400,0100) Items is not constrained.

#### A.34.xx.4.3 Number of Waveform Channels

The value of Number of Waveform Channels (003A,0005) in each Waveform Sequence (5400,0100) Item is not constrained.

#### A.34.xx.4.4 Sampling Frequency

The value of Sampling Frequency (003A,001A) in each Waveform Sequence Item is not constrained.

#### A.34.xx.4.5 Channel Source and Channel Source Modifiers

For Channel Source Sequence (003A,0208) in each Channel Definition Sequence (003A,0200) Item DCID 3001 “ECG Lead” shall​ be used.​

Note​

Terms from other Context Groups may also be used for extended specification of the Channel Source, as declared in the Conformance Statement for an application (see PS3.2).

#### A.34.xx.4.6 Waveform Sample Interpretation

The value of Waveform Sample Interpretation (5400,1006) in each Waveform Sequence Item shall be SS or SL.

#### A.34.xx.4.7 Waveform Annotation Module

For Concept Name Code Sequence (0040,A043) in the Waveform Annotation Sequence (0040,B020) DCID 3335 “ECG Annotation”​ shall be used.

# Changes to NEMA Standards Publications PS3.4Digital Imaging and Communications in Medicine (DICOM)Part 4: Service Class Specifications

Add new SOP Class to PS 3.4 Annex B tables

## B.5 Standard SOP classes

The SOP Classes in the Storage Service Class identify the Composite IODs to be stored. Table B.5-1 identifies Standard SOP Classes.

Table B.5-1. Standard SOP Classes

|  |  |  |  |
| --- | --- | --- | --- |
| SOP Class Name | SOP Class UID | IOD Specification(defined in PS 3.3) | Specialization |
| … | … | … |  |
| **High Resolution ECG Waveform Storage** | **1.2.840.10008.5.1.4.1.1.9.1.x** | **High Resolution ECG IOD** |  |
| … |  |  |  |

# Changes to NEMA Standards Publications PS3.6Digital Imaging and Communications in Medicine (DICOM)Part 6: Data Dictionary

Add new SOP Classes to PS3.6 Annex A Table A-1. UID Values

|  |  |  |  |  |
| --- | --- | --- | --- | --- |
|  UID Value | UID Name | UID Keyword | UID Type | Part |
| … | … |  | … | ... |
| **1.2.840.10008.5.1.4.1.1.9.1.x** | **High Resolution ECG Waveform Storage**  | **HighResolutionECGWaveformStorage** | **SOP Class** | **PS3.4** |
| … |  |  |  |  |

# Changes to NEMA Standards Publications PS 3.17Digital Imaging and Communications in Medicine (DICOM)Part 17: Explanatory Information

Adapt SSSS.1.5.1 Mapping of Polysomnographic Data to DICOM

### SSSS.1.5.1 Mapping of Polysomnographic Data to DICOM

Neurophysiology time series SOP Classes relevant to sleep studies are:​

…

Non-neurophysiologic time series or video SOP Classes relevant to sleep studies, are:​

General ECG Waveform Storage The General **~~Electrocardiogram (~~ECG~~)~~ Waveform** Storage SOP Class is used to store digitized electrical signals from the patient cardiac conduction system collected on the body surface, which **~~has~~** **have** been ​acquired by an ECG modality or by an ECG acquisition function within an imaging modality or a​ recording device.

**High Resolution ECG Waveform Storage The High Resolution ECG Waveform Storage SOP Class is used to store digitized electrical signals from the patient cardiac conduction system collected on the body surface, which have been acquired by an ECG modality or by an ECG acquisition function within an imaging modality or a​recording device**