## Slide 1: DICOM Educational Conference Bangkok, Thailand
OCTOBER 3-4, 2019
DICOM Enterprise Imaging
And WHOLE SLIDE IMAGING
David A. Clunie
PixelMed Publishing, LLC

## Slide 2: Disclosures
Editor of the DICOM Standard (NEMA Contract)
Owner of PixelMed Publishing, LLC
Consulting for BKMedical, Canfield, Carestream, Imago, MDDX (Bioclinica)
Supported by NIH U24CA180918 QIICR, NCI Leidos BOA 29XS219 Task Order #05

## Slide 3: Wide Variety of Images
Cardiology
Bronchoscopy
Gastrointestinal Endoscopy
Hematology
Pathology
Surgery
Nuclear Medicine
Dental
Radiology
Dermatology
Ophthalmology
Podiatry
Vascular
Urology
Nursing
Electrocardiography
Scanned Documents
Slide of VA VISTA 1995-7 from Kuzmak P, Dayhoff R

## Slide 4: Visible Light IODs and SOP Classes
VL Endoscopic Image (IOD and Storage SOP Class)
VL Microscopic Image
VL Slide-Coordinates Microscopic Image
VL Photographic Image
Video Endoscopic Image
Video Microscopic Image
Video Photographic Image
VL Whole Slide Microscopy Image

## Slide 5: Ophthalmic IODs and SOP Classes
Ophthalmic Photography 8 bit Image
Ophthalmic Photography 16 bit Image
Ophthalmic Tomography Image
Ophthalmic Refractive Measurements (Lensometry, Visual Acuity, …)
Ophthalmic Visual Field Static Perimetry Measurements
Ophthalmic Thickness Map
Wide Field Ophthalmic Photography Stereographic Projection Image
Wide Field Ophthalmic Photography 3D Coordinates Image
Ophthalmic Optical Coherence Tomography En Face Image
Ophthalmic Optical Coherence Tomography B-scan Volume Analysis

## Slide 8: Interoperability
“the ability of two or more systems or components to exchange information and to use the information that has been exchanged”
IEEE Standard Computer Dictionary: A Compilation of IEEE Standard Computer Glossaries. 1990

## Slide 12
Meyer-Ebrecht D. [Electronic Archival System for X-Rays Images - Work proposal for a research project in the
years 1974 and 1975] Elektronisches Archivierungssystem für Röntgenbilder – Arbeitsvorschlag für ein
Forschungsprojekt in den Jahren 1974 und 1975. Hamburg, Germany: Philips Research Labs; 1973 Oct.
1973

## Slide 13: DICOM and Radiology Modality
Modality
PACS
Storage

## Slide 14: DICOM and Radiology Modality
Modality
PACS
Storage
Commitment
Worklist
Completed

## Slide 15: DICOM and Slide Scanner
Slide Scanner
PACS
Storage

## Slide 16: DICOM and Slide Scanner
Slide Scanner
PACS
Storage
Commitment
Worklist
Completed

## Slide 17: DICOM Modality to PACS
Modality
Archive
Modality
Modality
Modality
PACS +/- RIS
Manager
Workstations
Standard Boundary

## Slide 18: DICOM WSI to PACS
Archive
PACS +/- APLIS
Manager
Workstations
Standard Boundary
Slide Scanner
Slide Scanner
Slide Scanner
Slide Scanner

## Slide 19: DICOM – Radiology Workstation
Modality
Archive
Modality
Modality
Modality
PACS +/- RIS
Manager
Workstations
Standard Boundary
Standard Boundary

## Slide 20: DICOM – Pathology Workstation
Archive
PACS +/- APLIS
Manager
Workstations
Standard Boundary
Standard Boundary
Slide Scanner
Slide Scanner
Slide Scanner
Slide Scanner

## Slide 21: DICOM – Analysis Systems
Archive
PACS +/- APLIS
Manager
Workstations
Standard Boundary
Standard Boundary
Slide Scanner
Slide Scanner
Slide Scanner
Slide Scanner
Analysis Systems

## Slide 22: DICOM – Enterprise Imaging
Archive
PACS +/- IS
Manager
Standard Boundary
Standard Boundary
Camera
Slide Scanner
Modality
Modality
Workstations
Analysis Systems

## Slide 23: DICOM – Deconstructed PACS
Archive
Manager
Standard Boundary
Standard Boundary
Camera
Slide Scanner
Modality
Modality
Workstations
Analysis Systems

## Slide 24: Cloud
Archive
Manager
Standard Boundary
Standard Boundary
Camera
Slide Scanner
Modality
Modality
Workstations
Analysis Systems

## Slide 25: Pathologist/Department
Archive
Manager
Camera
Slide Scanner
Modality
Modality
Workstations
Analysis Systems

## Slide 26: Enterprise IT (Security, Reliability)
Archive
Manager
Camera
Slide Scanner
Modality
Modality
Workstations
Analysis Systems

## Slide 27: Why DICOM?
Enormous experience in radiology and cardiology
34 years since ACR-NEMA PS3 Standard (1985)
A consensus of user and industry representatives; later adopted by ISO as ISO 12052
80 million CT studies per year in US (CBS News, 2015) – all DICOM
Huge supporting infra-structure – for both DICOM file format, protocol and services
All manner of products essentially commoditized: scanners, archives, workstations, viewers, PACS, toolkits for products, testing, analysis, research
Both commercial and free, closed and open source tools
Conformance and interoperability testing venues (e.g., IHE Connectathons)
Modality agnostic – e.g., XR, MR, NM also Visible Light, esp. Ophthalmology, Endoscopy
Application agnostic – human, veterinary, small animal research, non-destructive testing (esp. aerospace and nuclear power), security (esp. baggage scanning)
Emphasis on reliable, consistent, standard metadata (common data elements, value sets)

## Slide 28: Why not DICOM?
More effort than most trivial file formats – toolkits are generally required
Complexity is implicit in the use case more than the “format” per se – harder problems require more effort and discipline to be interoperable
Population of metadata takes effort – is it worth that effort?
Traditional DICOM network transport protocols are unique, though TCP/IP based – mitigated through more recent use of HTTP (WADO) using XML, JSON metadata
Pixel data encoding not a perfect match for WSI virtual microscopy – questions of size limits and tile access – multi-frame tiles are a hack (like TIFF), but are workable
Intellectual property (patent) distractions for WSI – now resolved
Legacy of use of proprietary (albeit mostly TIFF-based) – why change if downstream users/apps are willing to cope?
DICOM Conformance is not a panacea – claims of support are limited to query, storage and retrieval, worklists, etc., but NOT visualization (but DICOM does enable viewers)

## Slide 29: Status quo for WSI
Hodgepodge of proprietary file formats
Some (Big-)TIFF-based (good), some not (bad)
Some with extensions to TIFF (e.g., JPEG 2000 compression)
Some disclosed publicly, some not
Usually used with vendor-supplied viewer or proprietary SDK
Possibly readable by open source or 3rd party
Limited integration of scanners with Anatomical Laboratory Information Systems (APLIS), if at all, perhaps requiring expensive customization
No metadata: fragile linkage to contextual data (patient, slide, handling, staining) by filename or scanned slide identifier only
When decoupled from environment (APLIS, proprietary PACS), lose contextual data

## Slide 30: DICOM WSI – 2005 to 2019
1999 – Sup 15 – Visible Light including Microscopy
2005 – WG 26 got to work on WSI etc.
2006 – IHE Anatomic Pathology Domain
2008 – Sup 122 – Specimen Module
2008 – IHE Anatomic Pathology Workflow
2010 – Sup 145 – Whole Slide Microscopic Image IOD
… seven years of silence …
2017 – 1st premarket approval for primary diagnostic use
2017 – 1st WG 26 Digital Pathology Connectathon (PV)
2018 – three Connectathons (PathInfo, ECDP/NDP, PV)
2019 – 2nd premarket approval for primary diagnostic use
2019 – 6th Connectathon upcoming – 11 vendors registered

## Slide 31: DICOM WSI – What and How
File format for:
  whole slide images (tiled pyramid)
  single fields – slide microscopy
  gross microscopy
File contains:
  compressed pixels (JPEG or JPEG 2000)
  metadata – identifying and descriptive
Protocol for sending and receiving, etc.
Other stuff like workflow, annotation, segmentation, structured reports, …

## Slide 32
How digital slides are stored in a pyramid structure.
Wang Y, Williamson KE, Kelly PJ, James JA, Hamilton PW (2012) SurfaceSlide: A Multitouch Digital Pathology Platform. PLOS ONE 7(1): e30783. https://doi.org/10.1371/journal.pone.0030783
http://journals.plos.org/plosone/article?id=10.1371/journal.pone.0030783

## Slide 33
Garcia-Rojo et al. 2016

## Slide 34: DICOM WSI: Why tiled pyramids?
Goal is simplicity of access simulating a microscope
Zoom and pan
Tiles (frames): allow access to rectangular sub-regions of each resolution layer (without loading entire huge object)
Pyramid: entire highest resolution layer is very large, so storing lower magnification layers (for faster zooming) takes little extra space (about 30%)
Works around DICOM single frame size limitations (64k x 64k): no change to underlying DICOM encoding, no change to existing DICOM toolkits and archives
Do need services for metadata (index: which tile is which frame) and frame-level retrieval – WADO-RS

## Slide 35: Why reinvent TIFF?
TIFF (or BigTIFF) is the basis of many proprietary formats
Open Microscopy Environment (OME) TIFF is a good format
Data model and XML metadata
Microscopy-specific
Strong in the research community (esp. non-WSI microscopy)
Open source library support
Real question is: why reinvent DICOM when DICOM is used for all other images in the hospital?
TIFF usage -> just another silo (pathology), even if it is “open”

## Slide 36
https://www.teeturtle.com/products/have-your-cake-and-eat-it-too

## Slide 37: Dual Personality DICOM-TIFF
DICOM file format was designed to coexist with a second format
Bulk data (compressed pixels) shared between both formats
E.g., a single stored file can be both DICOM and OME-TIFF
Ideally would harmonize DICOM and OME metadata so that not only are pixels shared but identification and descriptions consistent
Mechanism: use 128 byte preamble to contain TIFF Image File Directory (IFD) to points to Dataset Trailing Padding after DICOM content, points back to payload of DICOM Pixel Data element
Both DICOM and TIFF use sufficiently similar JPEG encoding of pyramidal tiles to make this work for WSI

## Slide 40
http://medium.com/digital-trends-index/its-the-metadata-stupid-12a4fc121e45#.4zhwdz5y0
Notes: https://medium.com/digital-trends-index/its-the-metadata-stupid-12a4fc121e45#.4zhwdz5y0

## Slide 44: Specimen: Identification
Accession Number
  unit of work (case: order- report)
  same as radiology
“Specimens in Containers”
Specimen Identifier
  single discrete physical object considered a unit in workflow
Container Identifier
  part, cassette, block, section, slide, …
  container components … coverslip, etc.
No “Slide Identifier”, etc. per se – coded Container Type
Flexible: more than one specimen per slide (container), etc.

## Slide 45: Specimen: Description
Everything is coded with standard codes (mostly SNOMED CT)
Container Type Code Sequence = (433466003, SCT, "Microscope slide”)
Specimen Type Code Sequence = (430856003, SCT, "Tissue section”)
Primary Anatomic Structure Sequence = (10200004, SCT, "Liver")
(17636008, SCT, "Specimen Collection") = (86273004, SCT, "Biopsy ")
(111704, DCM, "Sampling Method") = (434472006, SCT, "Block sectioning")
(424361007, SCT, "Using substance") = (12710003, SCT, ”Hematoxylin stain")
(424361007, SCT, "Using substance") = (36879007, SCT, ”Water soluble eosin stain")
(430864009, SCT, "Tissue Fixative") = (431510009, SCT, "Formalin")
(430863003, SCT, "Embedding medium") = (311731000, SCT, "Paraffin wax")
Illumination Color Code Sequence = (414298005, SCT, "Full Spectrum”)
Illumination Type Code Sequence = (111744, DCM, "Brightfield illumination”)

## Slide 46: Anatomic Metadata

## Slide 47: Anatomic Metadata

## Slide 48: DICOM PACS
JPEG
JPEG Metadata – EXIF, ICC Profile
– Pixel Structure
Compressed Pixel Data Bitstream
DICOM Metadata – Identification – Acquisition Description
– Pixel Structure
Database Index of (subset of) Metadata
File system +/- metadata in filename or pathname
Camera
System
Viewer
DICOM Protocol

## Slide 49: EMR Export, Migration, Analysis, …
JPEG Metadata – EXIF, ICC Profile
– Pixel Structure
Compressed Pixel Data Bitstream
File system +/- metadata in filename or pathname
What patient?
What body part?
What encounter?
What date?
…

## Slide 50
Detachment
Sucks!
without embedded metadata, that is
https://imgur.com/gallery/AuD7M
Notes: https://imgur.com/gallery/AuD7M

## Slide 51: PV 2017 Connectathon
Archive
PACS +/- APLIS
Manager
Workstations
DICOM C-STORE
DICOM WADO-RS
Slide Scanner
Slide Scanner
Slide Scanner
Slide Scanner

## Slide 52: DICOM WG 26 WSI Connectathons Participation to date
 | PV’17 | PI’18 | ECDP’18 | PV’18 | ECDP’19
AidPath | View |  | Archive, View |  | 
Corista |  | Analyze |  |  | 
Gestalt |  |  |  | Archive, View | 
Neagen |  |  |  | Archive, View | Archive, View
PathCore | Archive, View | Archive, View |  | Archive, View | 
Sectra |  | View | View | View | 
3DHistech |  |  |  |  | Scan
Hamamatsu |  | Scan | Scan |  | Scan
Leica | Scan | Scan |  | Scan | 
Motic |  |  |  | Scan | 
Philips | Scan |  | Scan | Scan | Scan
Roche | Scan | Scan | Scan | Scan | Scan

## Slide 53: PV 2017 Connectathon Lessons
which compression schemes (JPEG, or J2K as well?)
one layer or entire pyramid from source (viewers expect latter, who makes it?)
how to recognize which pyramid layer is which (PixelSpacing)
recognizing a pyramid, in one series, multiple series, multiple per series
natural order of encoded frames versus their index
sparseness: entire tile array or selected sub-regions
tile frame size: same for each resolution layer (e.g., localizer non-square?)
dimensions described or not?
localizer with index, or not? in same or separate series?
concatenations: splitting huge files for transfer, requires reassembly on receipt
is a label image needed, does it need a barcode? shared between pyramids?
what optional metadata in image, in query (esp. specimen preparation)?
specific server services/sequencing for viewing (find vs. metadata retrieve)
WADO-RS – retrieve or retrieve rendered (multipart MIME burden)
color consistency – importance of viewer applying embedded ICC profile

## Slide 54: PV 2017 Connectathon Lessons
Need more Connectathons! Need more testing!
More specific profiling of requirements
  DICOM CPs to fix details, clarify ambiguities, optimize for common use-cases
  WG26 or IHE “profile”?
  clarify patterns of use for specific use cases
  make choices where alternatives exist, require currently optional features
Just works, or works for the right reasons?
  importance of validation against the formal standard requirements
  currently assisted by mechanical tools (dciodvfy) – could check more
  avoid using extensions, options, even if agreed upon
  check with proxy between devices (as used by IHE)
  create synthetic objects (good & bad) to stress recipients

## Slide 56: Beyond storage/retrieval …
Color management
  color normalization
  color consistency - ICC profiles
  services for application of ICC profiles to simplify (Internet browser based) viewers
Workflow management
  provision of identification and specimen preparation
Annotations
  input (“hot spots”) and output from analysis algorithms
  DICOM Segmentations
  DICOM Structured Reports
  ? something new in DICOM that scales to millions of nuclei, membranes, etc.

## Slide 57
No ICC Profile Applied
With ICC Profile Applied

## Slide 58: Standard Workflow Integration
Slide Scanner
PACS
DICOM
C-STORE
AP-LIS
DICOM
MWL
Standard Images and HL7/DICOM IS Integration
Broker
HL7
OML/QBP

## Slide 59: Standard Workflow Integration
Slide Scanner
PACS
AP-LIS
DICOM
C-STORE
Broker “improves” DICOM with IS Metadata
Broker
DICOM
C-STORE
HL7 QBP

## Slide 60: Camera wrapper application
PACS
DICOM
C-STORE
Wrapper provides user interface to populate metadata
Wrapper
Proprietary
Interface
Camera

## Slide 61: IS Integration for Metadata
PACS
IS
DICOM
C-STORE
Wrapper
DICOM MWL, HL7 QBP
Camera
Proprietary
Interface
Wrapper re-uses IS-supplied metadata

## Slide 62: DICOM Camera with MWL
PACS
DICOM
C-STORE
IS
DICOM
MWL
Standard Images and HL7/DICOM IS Integration
Broker
HL7
OML/QBP
Camera

## Slide 64
Yoon et al. Tumor Identification in Colorectal Histology Images
Using a Convolutional Neural Network. J Digit Imaging. 2018 Jul 31;1–10.

## Slide 65
Wen et al. A methodology for texture feature-based quality assessment
in nucleus segmentation of histopathology image. JPI. 2017.

## Slide 66: Conclusions
DICOM has had greatest success in radiology and cardiology
But there is a long tradition of supporting all forms of visible light imaging (esp. ,VA), now branded as “enterprise imaging”
DICOM provides not only image encoding, but robust, portable and interoperable, common and application-specific identifying and descriptive metadata, unlike any other format
DICOM provides protocols and services to support storage, query, retrieval and workflow
Challenging applications like Whole Slide Imaging are supported in DICOM by creative re-use of existing mechanisms (multi-frame images for tiled pyramids), which in turn allows re-use of existing archives and access protocols
Re-use of DICOM for other ‘ologies leverages other DICOM features, like ICC-based color management and annotation