## Slide 1
Supplement 223 Inventory IOD and Related Services
DICOM WG-33 Data Archive and Management

## Slide 2: History
WG-33 established by DSC December 2019
  In response to Laitek proposal from April 2019
Organizational meeting January 2020
  Elected officers: Matt Bishop, user co-chair; Keith Ecklund, vendor co-chair
Substantive work began March 2020, with broad user and vendor participation
  20+ participants per bi-weekly meeting typical
Consensus technical approach agreed August 2020
  Discussed ~24 major technical issues (see Closed Issues in Sup 223 draft)
Work Item approved by DSC August 2020
Draft Supplement approved for Public Comment by WG-6 March 25, 2021

## Slide 3: Primary Use Case – PACS migration
Users typically replace PACS after ~12-15 years, often with change of vendor
  Became a significant user (PACS administrator) concern in 2010’s
Institutional consolidation often requires merge of disparate PACS
Replacement/consolidation requires migrating historical data to a new system
  Digital storage often retained “forever”
  Volumes increasing with hi-res, 3D/4D, and multimodality studies
  Many institutions store > 109 instances with data volumes > 1015B (petabyte)
Performance at scale is challenging
  DICOM Query/Retrieve designed for accessing relatively small number of studies
  Need mechanisms that account for enterprise-scale data operations

## Slide 4: But Sup 223 is not a PACS migration standard
PACS migration is a multi-faceted endeavor, some aspects of which can be supported by DICOM standards, and others that cannot
Sup 223 addresses one critical need to support migration – production of a standardized inventory of the DICOM content of a repository
  Required to prepare for migration, and to validate completeness of migration
  Many other processes involved in execution of migration
Sup 223 is a standard for “Inventory IOD and related services”
Inventory object and services may also support other use cases

## Slide 5: Secondary Use Case – Clinical Research and Business Analytics
Inventory can be important data source (in conjunction with EMR and image instances) for research and analytics
  Representation of the PACS database in standard non-proprietary format
  Can be searched or processed without impacting PACS in clinical operation
  Transferrable to research IT system (data warehouse) – “extract” in paradigm of Extract/Transform/Load (E/T/L)
Inventory may be produced for a specified census of candidate patients, with further selection based on modality or study or series description
De-identification as required in research not in scope of Sup 223
  De-ID is part of E/T/L “transform”, and highly dependent on merge process with other data sources

## Slide 6: Secondary Use Case – Safety Backup
Image access is a patient safety / mission critical resource
PACS DBMSs typically have fault tolerant designs
  E.g., redundant online storage and offline backups
  But data is in proprietary format and dependent on the DBMS for effective use
DBMS becomes single point of failure
  May become inoperable, e.g., due to license key expiration or malware attack
Inventory is a DBMS-independent replica of the critical data content of the PACS database for the managed DICOM SOP Instances
If the repository instances are in DICOM File Format and referenced in the inventory, there is a complete alternate path to access the images

## Slide 7: Secondary Use Case – Quality Control
Inventory may be produced for studies missing critical attribute values
  Patient ID, Accession Number, Study Date
  Allows proactive correction
Comparison data for wellness check (continuous testing) of PACS operation
  C-FIND result for sample of studies compared to data in previous inventory (e,g., metadata, number of images)
  C-MOVE retrieve objects compared to direct filesystem access objects
  Message digest (MAC) computed on objects retrieved through filesystem compared to MAC in previous Inventory

## Slide 8: Overview of Sup 223

## Slide 9: Sup 223 focus - inventory object
Inventory of entire PACS archive (or described subset) in non-proprietary format
  IOD with hierarchical structure based on DICOM information models
  Encoded using DICOM formats and data structures
Inventory written to persistent objects (SOP Instance(s) / files)
  Accessible asynchronously
Functionally equivalent (+/-) to Query (C-FIND/QIDO) response for entire PACS database
  But Query typically has restrictions on number of responses
  Query requires synchronous network protocol response, infeasible for a billion records

## Slide 10
Archive
Data
Store
PACS
database
Inventory
Client
Inventory is representation of PACS DB
provided to client

## Slide 11
Inventory
references
1
1
Study
0-n
has component
1
Series
0-n
has component
1
Instance
0-n
has subject
has context
Patient
Imaging Service Request
1
1
1
1
Access mechanisms / links
Inventory Information Entity-Relation Model
Compare to Query Study Root Information Model
creates
Equipment
1
1-n

## Slide 12: Inventory IOD Structure
SOP Common attributes
Equipment attributes
Inventory attributes
Included Inventory Instance Sequence
Inventoried Studies Sequence
Inventory status attributes
>Study attributes
>Study data access attributes
>Imaging Service Request attributes
>Patient attributes
>Inventoried Series Sequence
>>Series attributes
>>Series data access attributes
>>Inventoried Instances Sequence
>>>Instance attributes
>>>Instance data access attributes
>Inventory SOP Class/Inst UIDs
>Inventory data access attributes
>Included Inventory Instance Sequence
>

## Slide 13: Inventoried Studies Sequence
Hierarchical structure using Sequence attributes, following Inventory Information Entity-Relation Model (Study Root)
Patient and Imaging Service Request attributes in Study level
Each level provides (optional) direct data access links (non-DICOM)
Inventoried Studies Sequence
>Study attributes
>Study data access attributes
>Imaging Service Request attributes
>Patient attributes
>Inventoried Series Sequence
>>Series attributes
>>Series data access attributes
>>Inventoried Instances Sequence
>>>Instance attributes
>>>Instance data access attributes

## Slide 14: Complex relationship of Study :: Imaging Service Request
Potentially complex relationship between the Study and Imaging Service Requests (ISR) in the real world
  See IHE RAD Scheduled Workflow Integration Profile (n MSPS :: m MPPS)
Inventory information model follows the basic DICOM Study information model supporting a single Accession Number representing an ISR
  See General Study Module PS3.3 Section C.7.2.1
If a Study has multiple associated ISRs, the request attributes are encoded at the Series level in the Request Attributes Sequence
Inventory IOD includes the Request Attributes Sequence to support this use

## Slide 15: Level of Inventory
Inventory may be produced at three levels - with Study records only, with Study and Series records, or Study, Series, and Instance records
Supports different approaches to migration, or non-migration uses
Production of Study-only Inventory expected to be significantly less resource-intensive than Instance-level Inventory

## Slide 16: Inventory describes access to stored instances
Data migration facilitated by direct filesystem access to stored SOP Instances
Functionally equivalent (+/-) to network Retrieve (C-MOVE/WADO)
  But Retrieve has performance issues for non-conventional access patterns and access through clinical operational pathways in PACS
Inventory specifies available access mechanisms
  Always default DIMSE (C-MOVE, C-GET) or DICOMweb
  Non-DICOM file access protocol (e.g., NFS, SMB, HTTPS, …) to target Part10 files, Part10 files in ZIP or TAR[+GZIP], and Part 10 files in Study- or Series-based folders
Inventory may record Message Authentication Code (cryptographic message digest) to ensure file integrity
  Reader may independently compute MAC and compare to value in inventory

## Slide 17
Archive
Data
Store
PACS
database
Inventory
Client
C-MOVE / C-GET
DICOMweb Studies Retrieve
File/Object Access
Inventory
Inventory describes access to archive data:
DIMSE, DICOMweb, non-DICOM protocol

## Slide 18: Non-DICOM Protocol File Access
Study
  [Base URI +] Folder URI all study files in folder
  [Base URI +] File URI (Container) all files in ZIP/TAR
  Series
  [Base URI +] Folder URI all series files in folder
  [Base URI +] File URI (Container) all files in ZIP/TAR
  Instance
  [Base URI +] File URI Pt10 file
  [Base URI +] File URI (Container) + Filename in Container Pt10 file in ZIP/TAR
Base URI = <scheme>://<authority>/[<commonpath>/]
  e.g., nfs://pacs.exampleinstitution.org/JZ0078555/
  Set at Study and/or Series level
  If used, Folder or File URI is relative path beginning with ./

## Slide 19: Accommodation for metadata updates
Concern - updates to metadata attributes may not (yet) have been propagated into stored SOP Instance files retrieved through file access protocol
  Recognition of common PACS implementation (descriptive, not prescriptive)
  Important technical mechanism supporting efficient archive operations when update of substantial numbers of objects is invoked, or objects have long access time
Inventory includes current metadata attributes that have not yet been propagated into stored SOP Instance files
  Client directly accessing files must always obtain current metadata from inventory

## Slide 20
Archive
Data Store [-updates]
PACS
database
[+updates]
Inventory[+updates]
Client
File/Object Access
Updated Metadata
Inventory provides current metadata for directly accessed stored objects

## Slide 21: Inventory may be for subset of complete archive
“Scope of Inventory” similar to C-FIND key attribute matching
  Minimal, broadly supported keys as required for migration use case
  Scope described in Inventory header
Studies in datetime range
  Study Date / Study Time
  Study Update DateTime (if present, otherwise defaults to Study Date/Time)
Studies with specific modalities, or study or series descriptions
Studies for list of Patient IDs [+Issuer], list of Study UIDs
Missing attribute values: Patient ID, Accession Number, Study Date
  Important for migration and study linkage to medical record
  No method in C-FIND to identify records with missing values

## Slide 22: Support for deprecated objects
Many PACS support deprecated objects (‘soft deleted’ or ‘hidden’)
  Including images rejected for quality or patient safety reasons through PACS UI or IHE Imaging Object Change Management (KO objects with specific title concept codes)
  But SOP Instances may be required to be archived for regulatory reasons (e.g., patient X-ray exposure)
Inventory includes Studies/Series/SOP Instances identified as Removed from Operational Use, with specific attribute/reason code
  ‘Operational use’ includes diagnostic, clinical, therapeutic, and administrative uses
  Such SOP Instances might not appear in normal C-FIND responses, but are included in inventory
Inventory includes IOCM KO objects if included in archive
  PACS simply reports what is has in archive
  Up to the receiving app (migration client) to determine what to do about them – different organizations may have different policies

## Slide 23: Multiple inventory objects
Enterprise inventories may be enormous
For implementation specific reasons content of an inventory may need to be split into more than one SOP Instance
  Practical limits on the maximum size of an individual SOP Instance - tractable object sizes
  Archive inventoried by multiple parallel processes
Logical inventory structured as tree of Inventory SOP Instances
  Inventoried studies in subsidiary nodes included by reference (Incorporated Inventory Instance Sequence)
  Root instance provides scope, final production status, and links to all instances
  More on this later

## Slide 24
Archive
Data
Store
PACS
database
Inventory object 1
Inventory object 2
Inventory may be divided into multiple instances organized in tree

## Slide 25: Associated Services

## Slide 26: Inventory as DICOM Non-Patient Object for Query / Retrieve / Store
Inventory defined within category of DICOM Non-Patient Objects
  Color Palette, Hanging Protocol, Defined Protocol, etc.
Simple addition to Part 4 Annex GG defines Storage Service (C-STORE transfer)
New Part 4 Annex for Inventory Query/Retrieve (C-FIND, C-MOVE, C-GET)
  Inventory Q/R Info Model
  Based on other NPO Q/R Services
Simple addition to Part 18 Section 12 defines equivalent DICOMweb operations
  Standard transcoding to XML and JSON for DICOMweb retrieve
Optionally stored in a Part 10 conformant file - transferred via Media Service, or accessed via non-DICOM file access protocol
  Similar to accessing Part 10 files from archive

## Slide 27
Inventories
Archive
Data
Store
PACS
database
Inventories
Client
C-FIND / C-MOVE / C-GET
DICOMweb Non-Patient Search / Retrieve
File/Object Access
Inventories
Inventory accessed by DIMSE, DICOMweb, non-DICOM protocol

## Slide 28: Network service to initiate inventory creation
App external to PACS/VNA may request production of inventory, especially for dynamically determined subsets of archive
  Migration, research, or QA clients
New service modeled on Storage Commitment (already implemented in PACS)
  Inventory Creation IOD for DIMSE-N services
  Initiate, Cancel, Pause, Resume: N-ACTION from SCU
  Progress indication, Completion: N-EVENT-REPORT push from SCP
  Asynchronous mechanism – supports extended period to create inventory
Periodic progress indication with reporting interval requested by SCU

## Slide 29
Archive
Data
Store
PACS
database
Inventories
Client
N-ACTION Initiate Inventory
N-EVENT REPORT Status
Client initiates inventory production through DIMSE service

## Slide 30: Inventory Production Statuses
PROCESSING
COMPLETE
FAILURE
CANCELED
PAUSED
Resume
Initiate
Cancel
Cancel
Some status changes may be due to N-ACTION requests (labeled arrows)
Pause

## Slide 31: Each defined service is separable from others
Inventory Creation service not required - PACS may implement creation of inventory from its administrative UI
Access to Inventory instances at discretion of PACS – could use DIMSE C-STORE (with or without C-MOVE or C-GET), DICOMweb, or non-DICOM file access protocol
ID and location of Inventory instances may be supported by C-FIND or DICOMweb, or may be done out of band by non-DICOM means (e.g., email notification of filename)
But there is no DICOM Conformance to just an IOD – at least one instance exchange service must be supported to assert DICOM Conformance

## Slide 32
Inventory Creation
Inventory Storage
Inventory MOVE
Inventory Web Search/Retrieve
Inventory FIND
Non-DICOM
File Access
Protocol
Inventory GET
Admin
UI
Inventory Media Storage
Inventory IODs and associated services
Inventory IOD
Inventory
Creation IOD
Inventory Q/R Model
DIMSE
Web
Media

## Slide 33: Summary of Sup 223
Inventory IOD
Specifies (if available) direct filesystem access to stored instances
Inventory has current metadata (possibly not propagated to stored instances)
Inventory may be produced for subsets of full archive
Supports instances not for operational use (including IOCM controls)
In category of DICOM Non-Patient Objects with typical network services (C-STORE, C-FIND, C-MOVE, C-GET, and DICOMweb equivalents)
Service to initiate inventory creation (with Inventory Creation IOD)

## Slide 34: Level and Scope of Inventory Examples

## Slide 35: Complete inventory
Zero Length Scope of Inventory Sequence indicates “universal match” – all studies, with just study records, or all instances
Inventory Content Level | STUDY
Scope of Inventory Sequence | 
Inventory Content Level | INSTANCE
Scope of Inventory Sequence | 

## Slide 36: Inventory for moving to deep archive
All studies with Study Date prior to and including 2010
Inventory Content Level | INSTANCE
Scope of Inventory Sequence | 
>Study DateTime Range | \2010
>Study Update DateTime Range | 
>Study Description | 
>Series Description | 
>Modalities in Study | 
>Study Instance UID List | 
>Patient List Sequence | 
>Attributes Missing Value | 

## Slide 37: Incremental inventory
All studies with Study Update DateTime on or after July 1, 2022
Inventory Content Level | INSTANCE
Scope of Inventory Sequence | 
>Study DateTime Range | 
>Study Update DateTime Range | 20220701\
>Study Description | 
>Series Description | 
>Modalities in Study | 
>Study Instance UID List | 
>Patient List Sequence | 
>Attributes Missing Value | 

## Slide 38: Quality Control inventory
Studies missing Study Date, Accession Number, or Patient ID
Inventory Content Level | STUDY
Scope of Inventory Sequence | 
>Study DateTime Range | 
>Study Update DateTime Range | 
>Study Description | 
>Series Description | 
>Modalities in Study | 
>Study Instance UID List | 
>Patient List Sequence | 
>Attributes Missing Value | 00080020 00080050 00100020

## Slide 39: Patient Safety Audit inventory
CT Studies with a Structured Report for patients in list externally generated; inventory may be further searched, e.g., for X-Ray Radiation Dose Report objects in the SR Series that will be retrieved
Inventory Content Level | INSTANCE
Scope of Inventory Sequence | 
>Study DateTime Range | 
>Study Update DateTime Range | 
>Study Description | 
>Series Description | 
>Modalities in Study | CT\SR
>Study Instance UID List | 
>Patient List Sequence | 
>>Issuer of Patient ID | 
>>Patient ID List | 817701\870812\009100\201764\...
>Attributes Missing Value | 

## Slide 40: Research inventory
MR-PET head Studies in 2019
Inventory Content Level | INSTANCE
Scope of Inventory Sequence | 
>Study DateTime Range | 2019\2019
>Study Update DateTime Range | 
>Study Description | *HEAD*
>Series Description | 
>Modalities in Study | MR\PT
>Study Instance UID List | 
>Patient List Sequence | 
>Attributes Missing Value | 

## Slide 41: Business analytics inventory
Echocardiography Studies since 2018 that were referred from a particular practice; inventory will be further searched for names of relevant doctors in Referring Physician attribute
Inventory Content Level | SERIES
Scope of Inventory Sequence | 
>Study DateTime Range | 2018\
>Study Update DateTime Range | 
>Study Description | 
>Series Description | *ECHO*
>Modalities in Study | US
>Study Instance UID List | 
>Patient List Sequence | 
>Attributes Missing Value | 

## Slide 42: Tree of Inventory Instances Included by Reference

## Slide 43: Inventory Content by Reference
Inventory object includes:
  Content Date/Time (Date/Time of inventory initiation)
  Equipment producing inventory
  Level and Scope of Inventory
  Links to other inventory objects included by reference, and/or
  Studies Inventory
  Completion status (with respect to Scope of Inventory at Content Date/Time on Equipment)
Linked objects form a tree from a root object
  Root specifies Content Date/Time, Scope of Inventory, and Completion status for tree
  Those attributes in non-root objects ignored (but are valid for their subtree)
Pre-existing inventory objects can be linked – may appear under multiple roots
Content Date/Time
Equipment
Level and Scope
other attributes
Incorporated Inventories Sequence
Inventoried Studies Sequence
Completion Status
Reference 1
Reference m
Study 1
Study n

## Slide 44: Tree of Instances Incorporated by Reference
Content Date/Time
Equipment
Level and Scope
other attributes
Incorporated Inventories Sequence
Inventoried Studies Sequence
Completion Status
Reference 1
Reference m
Study 1
Study n

## Slide 45: Example Approach: Serial Production
Content Date/Time
Level and Scope
other attributes
Incorp. Inventories Sequence
Inventoried Studies Sequence
Completion Status PARTIAL
Study 1
Study m
Content Date/Time
Level and Scope
other attributes
Incorp. Inventories Sequence
Inventoried Studies Sequence
Completion Status PARTIAL
Reference
Study m+1
Study n
Content Date/Time
Level and Scope
other attributes
Incorp. Inventories Sequence
Inventoried Studies Sequence
Completion Status COMPLETE
Reference
Study n+1
Study z
Objects filled with inventoried studies until max size, next object links to prior, last becomes root

## Slide 46: Example Approach: Baseline and incremental update
Content Date/Time <time 1>
Level and Scope
other attributes
Incorp. Inventories Sequence
Inventoried Studies Sequence
Completion Status COMPLETE
Study 1
Date <time 1>
Study m
Date <time 1>
Content Date/Time <time 2>
Level and Scope
other attributes
Incorp. Inventories Sequence
Inventoried Studies Sequence
Completion Status COMPLETE
Reference
Study m
Date <time 2>
Study z
Date <time 2>
Study n+1
Date <time 2>
Study n
Date <time 1>
Baseline included by reference in incremental update inventory, changed study simply has new (complete) record

## Slide 47: Example Approach: Parallel inventory (multiple subsystems)
Content Date/Time
Level and Scope
Equipment <subsys 1>
other attributes
Incorp. Inventories Seq
Inventoried Studies Seq
Comp Stat COMPLETE
Study 1.1
Study 1.n
Study 2.1
Study 2.m
Study p.1
Study p.z
Content Date/Time
Level and Scope
Equipment <archive>
other attributes
Incorp. Inventories Sequence
Inventoried Studies Sequence
Completion Status COMPLETE
Reference 1
Reference 2
Reference p
Content Date/Time
Level and Scope
Equipment <subsys 2>
other attributes
Incorp. Inventories Seq
Inventoried Studies Seq
Comp Stat COMPLETE
Content Date/Time
Level and Scope
Equipment <subsys p>
other attributes
Incorp. Inventories Seq
Inventoried Studies Seq
Comp Stat COMPLETE

## Slide 48: Arbitrary tree

## Slide 49: Empty inventory
Content Date/Time <date>
Equipment <equipment>
Scope <specified modality>
other attributes
Incorporated Inventories Sequence
Inventoried Studies Sequence
Number of Study Records <0>
Completion Status COMPLETE
Affirmative declaration that at Content Date/Time, identified equipment had no studies matching Scope

## Slide 50: Multiple records for study – reconciliation
Study may appear multiple times in tree of inventory SOP instances
  Parallel production of inventories with replicated storage, or with different series on different subsystems
  Sequential production (baseline + increment) with change to study content
Each Study record complete as known by the creator of that Inventory instance
Consumer required to reconcile
  No consolidation/reconciliation requirement on producer – simplify production
  Consumer (migration client) typically needs to do reconciliation anyway due to merge with EHR data, errors in DICOM data, or idiosyncratic/poor PACS implementation
Inventory records identify equipment and time stamp to support reconciliation

## Slide 51: Inventory Links
Challenge for inventory user to locate and retrieve all referenced inventories in the tree
  No requirement for producer to support Inventory Query/Retrieve as SCP
Inventory includes access info with each link to referenced Inventory instance
  DIMSE Inventory Query/Retrieve Service, DICOMweb Non-Patient Instance Service, or non-DICOM file access protocol
Incorporated Inventory Instance Sequence defined recursively
  Tree structure specifies all subsidiary Inventory instances and their links
Incorporated Inventory Instance Sequence from referenced instance copied into referencing instance

## Slide 52: Links with hierarchical structure of subtree links
Links structure from referenced object added in toto as subsidiary attribute
Inventory 3
Inc Inventory SQ
Inventory 1
Inventory 2
Inventory 1
Inventory 5
Inc Inventory SQ
Inventory 3
Inventory 6
Inc Inventory SQ
Inventory 2
Inventory 4
Inventory 4
Inventory 5
Inc Inventory SQ
Inventory 3
Inc Inventory SQ
Inventory 4
Inventory 1
Inventory 2
Inc Inventory SQ
Inventory 1
Inventory 2

## Slide 53: Inventory links in related services
Related Services need to reference an inventory
  Inventory Creation complete N-EVENT REPORT
  Inventory Query Model C-FIND
With tree of Inventory Instances, do we give entire tree, or just the root?
Mechanics of specifying recursive structure in DIMSE-N services is problematic
Approach is to identify only root inventory object in N-EVENT REPORT and C-FIND
  With full tree and access links in Included Inventory Instances SQ, app only needs to open one object (root)

## Slide 54: Study Update DateTime

## Slide 55: Study Update DateTime – a new request for PACS implementation
A highly desired capability is the ability to inventory Studies that were added or updated since the previous inventory
  Implement by subset match key for study update time - but ...
PACS generally do not record study update time (except perhaps in audit trails)
Sup 223 proposes new attribute for Study Update DateTime
  As with Number of Study Related Instances (0020,1208), defined at Study Level, but not in composite SOP Instances
Sup 223 addresses workaround for its absence
  Even if implemented in PACS, may have no assigned value for historical studies

## Slide 56: Study Update DateTime Range matching
Study Update DateTime range may be specified for studies to be included in inventory
If a Study does not have a defined Study Update DateTime available for matching, the datetime range is matched to Study Date and Time as a default
  Study Update DateTime may not be present because PACS doesn’t support it, or because it was empty for whatever reason
  Conceptually, Study Date and Time is possibly close to the initial Study Update DateTime
Study Update DateTime included as Type 2 Study level attribute in Inventory IOD
  May be empty if not supported

## Slide 57: Implications for implementation
Client needs to deal with uncertainties of datetime range matching to either to actual Study Update DateTime or fallback Study Date / Time
  E.g., may need to obtain inventories with overlapping datetime range and reconcile duplicates
  Sensitive to delays characteristic of workflow at site - e.g., cardiology studies sent to PACS after reading in department (1-2 days delay)
Miss of demographic updates after prior inventory might be mitigated by processing of update audit trail (e.g., HL7 ADT message archive)
  Beyond scope of DICOM
Future efforts (e.g., IHE) possible to encourage PACS implementation of Study Update DateTime upon patient/procedure update

## Slide 58: What’s Not in Sup 223 Normative Specs (but in Informative Annex)

## Slide 59: Security issues (1)
Services associated with inventories have no specifications regarding access control or transport security, but those are necessary for any real-world deployment, and are implementation- and site-specific
  Same as all DICOM network services
  Site must also address access control and transport security for non-DICOM protocol
  Storing access credentials in Inventory is NOT appropriate
Files (Inventory and stored SOP Instances) have inherent security vulnerabilities
  Part 10 header
  ZIP and TAR executable content

## Slide 60: Security issues (2)
No specifications regarding encryption of data at rest (e.g., SOP Instances in the repository, or Inventory objects)
  DICOM does specify methods for secure (encrypted) files, but issues such as key management and distribution again are implementation- and site-specific
  Storing keys in Inventory is NOT appropriate
Applications using Inventories need robust validation and controlled introduction to operational environment
Resource over-use by inventory application may result in denial of service to critical functions

## Slide 61: Inventory object lifecycle management
No specifications for explicit management of lifecycle of Inventory objects
  At discretion of system managing inventory objects
  No method specified for client to request deletion, except when canceling inventory production
  Creator may specify expiration datetime in Inventory Creation Complete notification, and in Inventory C-FIND response

## Slide 62: Incremental access to Inventory SOP Instance content
Inventory objects may be very large; DIMSE C-STORE transfers whole object only
Nothing comparable to frame-based retrieve specified in DIMSE services
DICOMweb has inherent byte range retrieval (optional support)
Direct network filesystem object retrievals may allow interactive retrieval (protocol typically integrated into operating system filesystem I/O)
Low information entropy means compression is effective; compressed whole object transfer may be feasible
  Deflated Transfer Syntax
  ZIP or GZIP of Inventory in Part 10 file

## Slide 63: Use of multiple AE Titles
A PACS may support multiple AE Titles to provide different “views” of the repository
  Different primary Patient ID based on caller’s organization
  Specific repository subsystem
  Studies with research consent

## Slide 64: Sup 223 Open Issues for Public Comment

## Slide 65: 1.A Scope of Inventory
Only the defined attributes of the Scope of Inventory Sequence may be used to select a subset of the repository for inventory
  Semantics of matching defined for each attribute, not a general matching approach (unlike C-FIND)
  Standard Extended conformance is not permitted to extend this Sequence attribute
  Permitted use of different AE Titles for different views of repository
Is a general key attribute matching capability needed? For what use cases?
Would that use be better satisfied by client-side (user) filtering of a broader inventory scoped only by the limited set of key attributes?
Would a receiver of an Inventory instance be able to properly interpret the scope of inventory for arbitrary sets of key attributes that it might not have in its data model?

## Slide 66: 1.B Matching mechanism different
The matching specification for the Scope of Inventory is tailored to the specific use case of an inventory for migration
  Multiple values in Modalities in Study matched to Modality attribute in each Series with logical AND
  Combination matching of Issuer of Patient ID and single value from List of Patient IDs to Patient ID attribute
  Matching to attributes missing values (Patient ID, Accession Number, Study Date)
The matching semantics introduce another attribute matching mechanism into DICOM after the C-FIND mechanism and the Attribute Value Constraint (a.k.a. “Selector” mechanism, PS3.3 Section 10.25). The Scope of Inventory mechanism, while similar to C-FIND, is tailored for its use cases. These types of matching are difficult or not possible with the existing C-FIND or Selector mechanisms. Is this new mechanism warranted?

## Slide 67: 1.C Matching on SOP Classes in Study
If a general matching mechanism is not defined, should matching on SOP Classes in Study be specified? What are the use cases?
Should matching on multiple values be specified? If so, what should be the logic relationship amongst those matching values (AND or OR)?
Or is selection by Modality sufficient for the use cases, with a subsequent client-side filtering for SOP Classes?

## Slide 68: 2.A Study / Series level folders or container files
Each SOP Instance in the repository can have one or more direct file access links
As an optimization to allow direct file migration operations on sets of files, a link may be provided to a folder or container file (ZIP, TAR) that has all the stored objects for a Study or Series
  No restriction on objects from other Studies/Series in folder / container
Will presence of “extraneous” objects affect use of Study/Series containers in migrations or other use cases?
Should links to Study or Series level folders or container files forbid instances of other Studies/Series in those targets?
  If forbidden, should non-DICOM files, such as reports be allowed?

## Slide 69: 2.B Non-DICOM file access protocols
Non-DICOM file access protocols are left open to implementation
  “Note: Conformance specification may be facilitated by reference to IANA-registered scheme”
Should specific protocols be prescribed normatively? If so, which ones? How should different HTTP-based cloud storage protocol APIs be differentiated (e.g., AWS vs. Azure)?

## Slide 70: 3.A De-identification
No service specifications for producing de-identified Inventory SOP Instances
Existing Part 15 Basic Application Level Confidentiality Profile is oriented toward classic image SOP Instances
  PHI attributes in the top-level data set
  Possibly ineffective for Inventory IOD with PHI within Sequence attributes
  Not clear that use cases for the Inventory are appropriate for applying that profile
Is there is a need for a de-identification profile for Inventory instances?

## Slide 71: 3.B User Authorization
A repository might limit disclosure or retrieval of SOP instances, studies, or patients following a variety of authorization policies,
  Often based on the user’s identity and/or instance attributes.
How should the Standard address the application of such rules during the production of an Inventory?
  E.g., depending on their permissions, different users of the Inventory Creation service might get different content for the same Scope of Inventory – is this acceptable?
  How should such different outcomes be identified in the Inventory SOP Instances?

## Slide 72: 3.C Confidentiality/privacy attributes
A repository might have confidentiality/privacy attributes on certain patients (e.g., VIP status).
Should the Inventory IOD include standardized confidentiality/privacy attributes in the Patient IE, and if so, what attributes?
  Or should such attributes be handled external to the DICOM IOD, e.g., in the EMR?

## Slide 73: 4.A Pause/Resume
The Inventory Creation service defines a “paused” status, e.g., to support a temporary shortage of resources, or other production block during a potentially lengthy (multi-day) inventory creation process
  Pause triggered by server application is presumably cleared by human intervention
In addition to server reasons for pausing production, N-ACTION operations are defined allowing a remote client initiating Inventory Creation to request pause/resume for operational reasons outside the knowledge of the server
Is this a necessary remote control for DICOM network operations, or should all control be exercised directly on the server (PACS) administrative interface?
What are the implications for server implementation for remote Pause/Resume controls?

## Slide 74: 4.B “Fresh” Inventory
A server may create an inventory by reusing (by reference) a baseline inventory and creating only the incremental inventory of studies that have been added or changed since that baseline (see slide)
Should a client be able to request a “fresh” inventory created from the database without reuse of existing inventories?
  Note that a server could refuse to accept a request for a fresh inventory based on authorization or resource constraints
  If available, would clients always request a “fresh” inventory?

## Slide 75: 4.C Tree of References
An inventory may be divided into a set of multiple SOP Instances linked in a tree structure (see slide)
Each SOP Instance includes a full tree of references to the subsidiary instances in the set, including access attributes for those instances (see slide)
Is requiring this complete tree an undue burden on the producer of an inventory? (see slide)
As access attributes may become “stale” and become broken links, is inclusion of those links appropriate?
  Inventory heavily relies on links for accessing SOP Instances in repository – PACS should know how to manage links

## Slide 76: 4.D Production Monitoring
The client that initiates Inventory Creation receives N-EVENT-REPORT status notifications on a regular basis
Is there a need for other clients to also receive such status notifications, e.g., for an inventory initiated locally at the PACS administrative user interface?
Should such a service include a client-initiated subscription capability? (Compare the Unified Procedure Step Subscribe/Unsubscribe)